QubeCL laser driver
user manual
ppqSense s.r.l
ppqSense Development Team
Revision 1.0 , 2024.05.15

---
QubeCL laser driver user manual
Contents
1
Warranty
5
2
Specifications
6
2.1
Communication interfaces
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
6
2.2
Absolute maximum ratings .
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
6
2.3
Electrical
Specifications
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
6
3
Typical
performance characteristics
10
4
Introduction
13
4.1
Safety consideration
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
13
4.2
Part List
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
13
4.3
Electrical
power supply .
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
13
4.4
Connectors and LED indicators functions
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
14
4.4.1
Power connector
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
16
4.4.2
TEC connector .
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
16
4.5
QubeCL_Control
software installation .
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
16
5
Basic operation:
current and temperature controls
18
5.1
Set-up
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
18
5.2
Operating with the Control
Software
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
19
5.2.1
Connecting to the QubeCL
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
20
5.2.2
QubeCL_Control
appearance
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
20
5.3
Operation
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
23
5.3.1
Temperature control
operation
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
24
5.3.2
Current control
operation
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
25
5.3.3
Automatic protections
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
26
6
Advanced Operations - Modulators and DDS
28
6.1
The DDS sub-module
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
29
6.1.1
Connectors .
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
29
6.1.2
Control
Software interface .
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
29
6.1.3
DDS operations
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
31
7
Advanced Operations - PLL module
32
7.1
Connectors
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
32
7.2
Control
Software Interface .
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
32
7.3
PLL Operation .
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
34
8
Advanced Operations - PDH module
36
8.1
Connectors
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
36
8.2
Control
Software Interface .
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
36
8.3
PDH operations
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
38
Revision 1.0
Pag.
1 - 69

---
9
Advanced Operations - LIA module
40
9.1
Connectors
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
40
9.2
The DDS sub-module operation in pair with the LIA module .
.
.
.
.
.
.
.
.
.
40
9.3
Control
Software Interface .
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
42
9.4
LIA operations
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
43
9.5
LIA digital
filter
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
44
10 Advanced Operations - Slow drift compensation
46
11 Advanced Operations - QubeCL advanced settings
48
11.1 Adv.
1 tab .
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
48
11.2 Adv.
2 tab .
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
49
12 Interfacing with lasers
51
12.1 Laser Housing for QubeCL .
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
51
12.1.1 Butterfly Packaged laser Housing
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
51
12.1.2 HHL Packaged laser Housing
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
51
12.2 Laser Interface Boards
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
51
12.2.1 HHL Interface Board .
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
52
12.2.2 NanoPlus Cubic Housing Interface Board .
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
52
12.3 High Frequency modulation with Bias-Tee
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
52
12.3.1 Bias-Tee usage for high freuqency laser modulation .
.
.
.
.
.
.
.
.
.
.
52
12.3.2 Interface Boards Wiring
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
55
13 Direct Communication
56
13.1 QubeCL Serial
Port Specifications .
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
56
13.2 QubeCL Serial
Commands .
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
56
14 Customer support and useful
contacts
68
2 - 69

---
List of Tables
1
Communication Interfaces
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
6
2
Absolute Maximum Ratings
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
6
3
Current Generator Specs .
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
6
4
Analog Current Modulators Specs .
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
7
5
Mixed Current Generator Specs - Digital
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
7
6
Mixed Current Generator Specs - Analog
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
8
7
Temperature Controller Specs
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
8
8
PLL Module Specs .
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
8
9
Pound-Drever-Hall
module Specs
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
8
10
Lock-In Amplifier module Specs
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
9
11
Power source voltages specifications .
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
14
13
Available commands for QubeCL current generator.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
57
14
Available commands for QubeCL temperature controller.
.
.
.
.
.
.
.
.
.
.
.
.
58
15
Available commands for the DDS submodule.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
60
16
Available commands for the PLL locking module.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
62
17
Available commands for the PDH locking module.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
64
18
Available commands for the LIA locking module.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
66
19
Available commands to check the instrument status.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
67
List of Figures
1
Current noise spectral
density comparison with low current modulator
.
.
.
.
.
10
2
Current noise spectral
density comparison with high current modulator
.
.
.
.
10
3
24h Current stability over ambient temperature .
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
11
4
24h Current stability over time
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
11
5
Current temperature coefficient
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
12
6
QubeCL15TP front panel
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
15
7
Power connector pin diagram .
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
16
8
TEC connector pin diagram .
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
16
9
QubeCL control
software main screen .
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
19
10
The QubeCL_Control
software configuration panel
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
20
11
The QubeCL_Control
software current driver panel
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
21
12
The QubeCL_Control
software temperature control
panel
.
.
.
.
.
.
.
.
.
.
.
22
13
The QubeCL_Control
software Advanced tab .
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
23
14
Appearance of the LH module equipped with the DDS sub-module
.
.
.
.
.
.
29
15
The DDS control
panel
of the QubeCL_Control
Software
.
.
.
.
.
.
.
.
.
.
.
30
16
Appearance of the PLL module
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
32
17
Control
Software Interface dedicated to the PLL Module .
.
.
.
.
.
.
.
.
.
.
.
33
18
Appearance of the PDH module .
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
36
19
Control
Software Interface dedicated to the PDH Module
.
.
.
.
.
.
.
.
.
.
.
37
20
Appearance of the LIA module
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
40
21
Control
Software Interface dedicated to the LIA Module
.
.
.
.
.
.
.
.
.
.
.
.
42
3 - 69

---
22
The first Advanced tab of the Control
Software GUI
.
.
.
.
.
.
.
.
.
.
.
.
.
.
48
23
The first Advanced tab of the Control
Software GUI
.
.
.
.
.
.
.
.
.
.
.
.
.
.
50
24
Bias Tee simplified electronic diagram .
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
52
25
Drawings of the NanoPlus Cubic Mount Interface Board .
.
.
.
.
.
.
.
.
.
.
.
55
26
Drawings of the HHL Interface Board .
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
55
4 - 69

---
1
Warranty
ppqSense S.r.l.
guarantees its QubeCL System to be free of
material
and workmanship
defects for one year from the date of the original
shipment (1-YEAR WARRANTY).
This
warranty is in lieu of all
other guarantees expressed or implied, including any implied warranty
of merchantability or fitness for any particular purpose.
ppqSense S.r.l.
shall
not be liable for
any special, incidental
or consequential
loss.
During the warranty period, ppqSense S.r.l.
will
repair or replace the unit or module, or issue credit, at our option, without charge.
ppqSenses
liability shall not in any case exceed the cost of correcting defects in the products as explained
here, and this service is the sole remedy of the buyer.
This warranty does not apply to defects
caused by abuse, accident, modifications, Acts of God, or the use of the product for which it
was not intended.
ppqSense S.r.l.
shall not be liable for damages resulting from the use of the
product,
nor shall
be responsible for any failure in the performance of other items to which
the purchased product is connected or the operation of any system of which the purchased
product may be a part.
The QubeCL System should not be used in the manner not specified
by the manufacturer.
5 - 69

---
2
Specifications
2.1
Communication interfaces
Interface
Mode
Settings
Notes
USB
Virtual
Serial
Port
115200 8N1
WiFi
Access Point/Client
software configurable
Optional
Table 1:
Communication Interfaces
2.2
Absolute maximum ratings
Parameter
Values
Unit
Notes
◦
Operating temperature
-20 to +40
C
◦
Storage temperature
-20 to +70
C
Main power supply voltage (PS
)
26
V
No terminal
must be
LAS
TEC power supply voltage (PS
)
24
V
connected to ground
TEC
Modulation inputs
±10
V
Table 2:
Absolute Maximum Ratings
2.3
Electrical
Specifications
◦
PS
=24 V
,
PS
=12 V
,
Warm-up time 120 min,
room temperature 25
C,
unless
LAS
dc
TEC
dc
otherwise noted.
Parameter
Min
Typ
Max
Unit
Notes
√
Noise spectral
density
100
200
500
pA/
Hz
test current 200mA
Integrated noise
50
100
200
nA
10 Hz - 100 kHz
RMS
Integrated noise
250
350
500
nA
10 Hz - 1 MHz
RMS
Stability (1h)
4
5
6
ppm
FS
Stability (12h)
5
6
7
ppm
FS
◦
Temperature coefficient
-2
-3
-5
ppm
/
C
FS
Voltage compliance
17.5
V
Table 3:
Current Generator Specs
6 - 69

---
Parameter
Min
Typ
Max
Unit
Notes
Analog input voltage
0
±10
V
DC offset
20
30
50
µA
input shorted
Low current
√
Noise spect.
dens.
floor
50
100
pA/
Hz
input shorted
Output current
0
±5
mA
Gain
-0.5
mA/V
Modulation BW (-3 dB)
2
MHz
High current
√
Noise spect.
dens.
floor
200
2000
pA/
Hz
input shorted
Output current
0
±100
mA
Gain
-10
mA/V
Modulation BW (-3 dB)
600
kHz
Table 4:
Analog Current Modulators Specs
Parameter
Min
Typ
Max
Unit
Notes
Waveforms
Sine or Triangle
−3
6
Freq.
range
10
10
Hz
Min freq.
step
250
mHz
Amplitude control
res.
12
bits
DC offset
20
30
50
µA
input shorted
Low current
√
Noise spect.
dens.
floor
50
100
pA/
Hz
input shorted
Output current
0
±2.5
mA
Modulation BW (-3 dB)
2
MHz
High current
√
Noise spect.
dens.
floor
200
2000
pA/
Hz
input shorted
Output current
0
±75
mA
Modulation BW (-3 dB)
600
kHz
Table 5:
Mixed Current Generator Specs - Digital
7 - 69

---
Parameter
Min
Typ
Max
Unit
Notes
Analog input voltage
0
±2.5
V
DC offset
20
30
50
µA
input shorted
Low current
√
Noise spect.
dens.
floor
50
100
pA/
Hz
input shorted
Output current
0
±2.5
mA
Gain
-1
mA/V
Modulation BW (-3 dB)
2
MHz
High current
√
Noise spect.
dens.
floor
200
2000
pA/
Hz
input shorted
Output current
0
±75
mA
Gain
-30
mA/V
Modulation BW (-3 dB)
600
kHz
Table 6:
Mixed Current Generator Specs - Analog
Parameter
Min
Typ
Max
Unit
Notes
Thermistor
1
10
kΩ
NTC
TEC current
±3
A
TEC compl.
voltage
20
V
Temp stability
40
400
µK
absolute stability in 12h
Temp.
coefficient
-90
-100
µK/K
Table 7:
Temperature Controller Specs
Parameter
Min
Typ
Max
Unit
Notes
RF freq.
input
10
250
MHz
LO freq.
input
10
100
MHz
Input level
-30
0
dBm
Input stage gain
0
30
dB
Lock bandwidth
800
kHz
Table 8:
PLL Module Specs
Parameter
Min
Typ
Max
Unit
Notes
RF freq.
input
1
100
MHz
LO freq.
input
1
100
MHz
RF Input level
-30
0
dBm
LO Input level
5
dBm
Lock bandwidth
800
kHz
Table 9:
Pound-Drever-Hall
module Specs
8 - 69

---
Parameter
Min
Typ
Max
Unit
Notes
Demodulation frequency
32768
Hz
fixed
Ramp frequency
1
1000
Hz
Input voltage noise
30
µV
RMS
Input stage gain
0
60
dB
Integration time
1
10000
ms
Lock bandwidth
1
kHz
Table 10:
Lock-In Amplifier module Specs
9 - 69

---
3
Typical
performance characteristics
3
10
M2
2
10
M1
1
10
0
10
1
10
4
5
2
3
6
10
10
10
10
10
Frequency (Hz)
M1 = 50nA
M2 = 440nA
rms
rms
Figure 1:
Current noise spectral
density comparison with low current modulator
4
10
M2
3
10
2
10
M1
1
10
0
10
1
10
4
5
2
3
6
10
10
10
10
10
Frequency (Hz)
M1 = 50nA
M2 = 2µA
rms
rms
Figure 2:
Current noise spectral
density comparison with high current modulator
10 - 69
Current Noise PSD (nA/Hz
1/2
)
Current Noise PSD (nA/Hz
1/2
)

---
Figure 3:
24h Current stability over ambient temperature
Figure 4:
24h Current stability over time
11 - 69

---
Figure 5:
Current temperature coefficient
12 - 69

---
4
Introduction
This manual contains information for operating the modular QubeCL system.
The QubeCL
is a high-performance instrument that provides low-noise and high-ratings electric current and
temperature control
to drive optoelectronic devices.
The QubeCL is particularly suited to
drive state of the art semiconductor and Quantum Cascade laser (QCL) sources.
This chapter
addresses pertinent safety issues and correct usage of
this instrument,
and defines all
front
panel connections and LED function indicators.
Please, carefully read this chapter before using
the QubeCL.
4.1
Safety consideration
The QubeCL is a versatile instrument that can be used in a variety of driving current and
temperature conditions.
However,
the QubeCL is not intended for
a fail-safe operation in
hazardous environments or
life-threatening situations.
The user
assumes full
responsibility
for correct and safe usage of the QubeCL in accordance with any applicable laws, codes and
regulations, and standard pertaining to their specific application.
ppqSense S.r.l.
is not liable
for any consequential
damage due to misapplication or failure of the QubeCL system.
The QubeCL System is compliant with the following standards:
EN 61326-1, EN 55011, EN 61000-4-2, EN61000-4-8, EN61000-4-3
4.2
Part List
The QubeCL system is shipped in a package designed to provide excellent protection.
The
shipping box should be saved for future transportation or storage.
Carefully unpack and inspect
the following items that are contained in the shipping box.
• QubeCL driver,
built with a variable number and kind of modules depending on the
ordered model.
• Power supply cable.
• TC cable, to connect the temperature controller (only included if the QubeCL mounts
a Temperature Controller module).
• USB cable, to connect the QubeCL driver to the controlling PC.
• USB drive, which contains the required software and drivers.
• QubePS, switching power supply for the QubeCL driver, if included in the order.
4.3
Electrical
power supply
The QubeCL system must be powered by a unipolar
floating DC power
supply.
If
the
QubePS is not included in the QubeCL kit, a dual channel power supply or two different power
supplies are necessary to independently drive the Current Generator [CM] and the Temperature
Controller [TC] modules.
CM modules must be powered by a DC power supply providing 24 V and at least 3 A for
a versatile usage and for a full
compliance to the declared performances and dynamic range.
13 - 69

---
TC modules must be powered by a DC power supply providing 12 V and 4 A, for a versatile
usage and to guarantee performances of the TC module.
User must refer to Table 11 for minimum,
maximum and typical
allowable voltages for
power supply of the QubeCL. Please, note that using source voltages that differ from the ones
indicated as typical, even if in between the maximum and minimum values, may lead to poor
calibration of the QubeCL driver or to potentially damaging heat generation.
DC Input
minimum typical
maximum
LAS
+20 V
+24 V
+30 V
TEC
+8 V
+12 V
+24 V
Table 11:
Power source voltages specifications
CAUTION
Power supplies used to source the QubeCL drivers must be floating power supplies.
The terminals of the power supply must not be connected to Earth or to one another
in any circumstance.
If the QubePS power supply is included in the QubeCL kit, the user has simply to connect
the QubePS to the QubeCL with the dedicated cable delivered within the box.
The power
cord needed to connect the QubePS to the electrical
outlet is not provided.
4.4
Connectors and LED indicators functions
A schematic view of the front panel
of the QubeCL can be seen in figure 6.
The image
represents a QubeCL15TL,
being a QubeCL equipped with a 1.5 A current generator,
Tem-
perature Controller module and Phase-Locked Loop module.
Below,
a list of the connectors
and LEDs present on the front panel
can be found, along with their description:
Connectors:
• C1.
Power:
power supply connector of the instrument.
• C2.
USB: this USB port must be used to connect the QubeCL to an host computer in
order to control
and supervise its operations.
• C3.
TEMP out:
used to connect the supplied TC cable for the temperature control
of the driven device.
It provides NTC (10 k) temperature sensor input and TEC output
(2.7A maximum).
• C4.
CURRENT out:
Output SMA connector.
The driver supplies the current through
this connector to the controlled laser device.
A proper SMA cable must be provided by
the user.
The output provides a maximum current corresponding to the limits of the
installed CM modules of the QubeCL (for example 1.5A for a QubeCL with one CM10
and one CM05 modules), with a maximum output voltage compliance of 17.5 V.
14 - 69

---
• C5/C6.
MOD1/2 in:
Input SMA connectors.
May be used to source finely controlled
modulating signals to the laser current, with the amplitude proportional
to the applied
control
voltages.
Figure 6:
QubeCL15TP front panel
LED indicators:
• L1.
CONNECT: Green.
Switched on when the QubeCL driver is connected to an host
PC through its USB port.
• L2.
CUR ON: Yellow.
Switched on when the QubeCL sources current from C4 Current
out connector to the controlled laser.
• L3.
MOD ON: Yellow.
Switched on when the modulation signals are enabled.
• L4.
PWR ON: Green.
Switched on when the QubeCL main unit is powered on.
• L5.
TC unlock:
Red.
Switched on when the temperature of the controlled device is
different with respect to the temperature set-point value.
• L6.
TC lock:
Green.
Switched on when the temperature of the controlled device is
locked and equal
to the temperature set-point value.
The list only comprehends the LEDs and connectors that are present on the majority of
the QubeCL possible configurations.
A detailed description of PLL module, along with other
application-specific modules, can be found further below in this manual.
15 - 69

---
4.4.1
Power connector
The
QubeCL kit
comes
with a
dedicated power supply cable to con-
nect the QubeCL to its power supply.
One
of
the
ends
of
the
QubeCL
power cable is already equipped with
a 10 poles connector compatible with
the
one
present
on
the
QubeCL
driver (see figure 7).
If
the QubePS is provided with the
kit,
the cable is ready to be directly
Figure 7:
Power connector pin diagram
plugged into the QubePS connector.
Alternatively,
when the QubePS is not provided,
the cable is equipped with four plugs to be
connected to the power supplies.
The four plugs can be identified as the subsequent ones:
• two for CM power supply (green-labelled PS Las +, white-labelled PS Las -)
• two for TC power supply (red-labelled PS TEC +, yellow-labelled PS TEC -).
4.4.2
TEC connector
If
the QubeCL is equipped with the TC
module,
but not with a laser-housing mod-
ule,
a proper
connection with the Peltier-
based stage of the laser and with the NTC
temperature sensor must be established.
The QubeCL kit includes a TEC cable with
one of
its ends already equipped with the
proper connector to be plugged into the TC
module.
The user must adapt the other end of
the
TEC cable in order to connect it to the laser.
Figure 8:
TEC connector pin diagram
Please, refer to the wires colors represented
in figure 8 to properly connect the TEC ca-
ble to the laser.
4.5
QubeCL_Control
software installation
Every QubeCL system,
regardless of the specific model,
can be controlled by the mean of
the same Control
Software,
which comes in the USB key included in the QubeCL kit.
The
QubeCL_Control
software has been developed under LabView Runtime Environment.
To install the QubeCL_Control
software on Windows OS, please follow these steps.
All
the
needed softwares are provided inside the USB key that comes inside the QubeCL kit.
16 - 69

---
• Run the library installation application to install
the following libraries and drivers:
– Microcontroller USB-driver (CDM21228_Setup)
– LabVIEW NI-VISA Utilities (NIVISA1600full)
• Copy the directory "/QubeCL_CONTROL:vX.X " and all
its content in a directory of
your choice in the computer.
More information can also be found in the readme.txt file in the main directory of the USB
key.
17 - 69

---
5
Basic operation:
current and temperature controls
This chapter explains the typical
set-up and operation of a QubeCL system composed by a
CM module (or a stack of several
of them) and a TC module.
This chapter should help less
experienced users to rapidly adapt the QubeCL system to drive their specific optoelectronic
devices.
Keep in mind that not all the possible operations of the QubeCL system with a specific
controlled device can be adequately addressed in this manual:
these are general
guidelines for
operating a QubeCL system with specific illustrative examples.
Adequate parameter values
must be considered by the users when controlling a specific device.
The QubeCL is built with a modular approach, with each module being in charge of performing
a specific function.
This chapter describes the usage of a QubeCL which includes one CM and
one TC module.
The functions of those modules are, respectively:
• To provide a very low-noise and high-level
DC current to drive one optoelectronic
device;
• To provide the optoelectronic device temperature stabilization at a few mK level.
5.1
Set-up
The typical set-up to operate with a QubeCL system is composed by the QubeCL itself, the
laser device, a DC power supply and a PC for overall
control.
To correctly operate with the QubeCL, the following step must be followed:
• If the QubeCL kit includes a QubePS power supply, simply connect the QubeCL to the
QubePS with the dedicated cable available in the kit.
Otherwise, if the QubePS is not included:
– Set the voltage values of
the floating power supplies you’re using following the
indication from the dedicated section of this manual
(chapter 4.3);
– Connect the QubeCL to the power supply by the mean of the provided cable
∗ Connect
TEC+ and TEC-
to the +/-
poles of
the power
supply selected
channel
for electrical
supply of the TC module, respectively.
∗ Connect LAS+ and LAS- to the +/- poles of the power supply selected channel
for electrical
supply of the CM module(s), respectively.
CAUTION
If not using a QubePS be sure to RESPECT THE POLARITY of the power supply
cable.
REVERSE voltages will
damage the instrument.
In case you’re not using a dual-channel
power supply, but two different power supplies,
please switch on the TC supply before the CM supply otherwise the QubeCL won’t boot
correctly.
• Connect the provided power supply cable to the C1 connection of the main unit module
of the QubeCL system;
18 - 69

---
• Connect the provided USB cable to the C2 connection of the main unit module of the
QubeCL system;
• Connect the provided TC cable, properly wired to the laser, to the C3 connector of the
TC module of the QubeCL system.
N.B. This operation is not necessary if the QubeCL includes a laser housing module;
• Connect the user provided laser device current cable to the SMA C4 connector of the
LH module of the QubeCL system.
N.B. This operation is not necessary if the QubeCL includes a laser housing module;
• Optional:
connect a signal generator(s) to the modulation inputs of the QubeCL (SMA
C5-C6 connectors of the LC module).
5.2
Operating with the Control
Software
In this section the usage of the Control Software to manage the basic operation of the QubeCL
is described in detail.
Figure 9:
QubeCL control
software main screen
19 - 69

---
5.2.1
Connecting to the QubeCL
To use the QubeCL_Control
software,
locate the directory where you’ve previously placed
the QubeCL_CONTROL:vX.X folder from the USB key and launch the QubeCL_Control_vX.X.exe
application.
A graphical interface as the one showed in figure 9 will be displayed.
While the program search
for connected QubeCLs to the computer,
a progress bar will
be displayed.
Once the search
is done,
it will
be possible to select the desired QubeCL from the INSTRUMENT ADDR
drop-down menu.
Once selected,
launch the program by pressing the CONNECT button.
If the program fails to detect any connected QubeCL,
an error
windows as the one shown here pops up.
From there it will
be possible to stop the program
with the EXIT button or manually select a connected
driver by the COM port to which it is connected.
5.2.2
QubeCL_Control
appearance
The GUI of the control
software is subdivided into three main horizontal
panels:
Figure 10:
The QubeCL_Control
software configuration panel
The Configuration panel at the top (Figure 10), there can be found:
• INSTRUMENT menu:
allows the selection of
the QubeCL that the user
wants to
connect to.
The connection can be established only at the start-up of the program,
it
is not possible to switch to another QubeCL instrument without restarting the control
software.
• CONNECT and DISCONNECT buttons:
allow the connection and disconnection to
the selected QubeCL. DISCONNECT also stops the program.
• CURRENT LIMIT menu:
by pressing this button,
another window will
be displayed,
allowing the user to change the maximum limit for the current that the QubeCL can
deliver to the laser.
• System information menu:
shows some information about the QubeCL to which the
software is connected, such as firmware version, instrument model
and serial
number.
• Power ON indicators:
above the CONNECT and DISCONNECT buttons,
two LEDs
show the status of the power source to the CM (called MAIN) and TC modules.
20 - 69

---
The Current driver panel is the middle one (Figure 11), there can be found:
• Current setpoint:
from there, the user can change the current that needs to be sourced
by the QubeCL to the laser.
Three different controls allow the user to change this value:
– A numerical
value expressed in mA can be directly typed into the dedicated form;
– A horizontal
slider,
to the left of
the numerical
value,
can be used to set the
current;
– Four buttons can be used to increment or decrement the current by fixed steps.
The desired amplitude for the steps can be typed in the two dedicated forms below
the arrow buttons.
• Iout button:
enables/disables the current output of the QubeCL.
This button includes a green LED indicator.
This LED is switched ON when the QubeCL
is delivering current to the laser.
Once the user switches off the current,
the button
becomes gray and the QubeCL gradually lowers the current sourced to the laser until
it
is completely switched off.
The LED will
be switched off only when the current is.
• Mod En button:
enables/disables all
the modulation channels;
• M1 and M2 buttons:
activate/deactivate the modulation channels built on board of
the QubeCL. By enabling those channels, modulating currents can be injected over the
bias current;
• Save button:
saves the parameters currently set in this panel
to the ROM memory of
the QubeCL. Those values will
be recalled from the memory at the next startup of the
QubeCL driver;
• Readings window:
shows the values of the output current (Ilas) and voltage (Vlas) as
well
as the supply voltage (Vcc);
• Current plot:
shows the temporal
behaviors of the set current (orange trace) and the
actual
output current (purple trace).
This plot won’t show any modulation signal
that
may be applied to the output current.
Figure 11:
The QubeCL_Control
software current driver panel
21 - 69

---
The Temperature control panel at the bottom (Figure 12), where the subsequent com-
mands can be found:
• Temp.
SetPoint:
sets the desired operating temperature for the laser,
expressed in
°C.
• Temp.
Read:
returns the value of
the temperature measured on the laser using its
NTC sensor, expressed in °C.
• TStab button:
turns on/off the temperature stabilization of the laser.
• Loop sign +/- button:
allows the user to change the sign of the temperature stabi-
lization loop, effectively switching its behavior from heating to cooling and vice versa.
• Temperature PID parameters panel:
the TEC module stabilizes the temperature of
the laser by the mean of a PID correction loop.
This window allows the user to change
the values of the three functional
parameters of the PID, both with sliders and textual
inputs.
Detailed instruction can be found below.
• Save button:
saves the setpoint,
loop sign,
PID gains and Tstab status on the ROM
memory of the QubeCL so that they can be used as default at next switch on.
N.B. This command also saves the status of Tstab, hence if the Save button is pressed
when the temperature stabilization is active,
the QubeCL will
automatically start to
stabilize the temperature once it gets powered up.
If this behavior is unwanted,
Save
button must be pressed when the temperature stabilization is inactive.
• Temperature and TEC current plots:
switching from one to another,
the user can
have a look over the temporal
behavior of the readed and setpoint temperatures of the
laser and of the current erugated by the TC module.
Figure 12:
The QubeCL_Control
software temperature control
panel
22 - 69

---
5.3
Operation
Once the QubeCL has been correctly set-up,
the effective operations can begin following
the steps listed below:
• Turn on the floating power supply (or the QubePS if it is comprised in the kit) to supply
the instrument.
LED L4 (green) lights on.
LED L5 (red) lights on,
meaning that the laser temperature is not stabilized.
If
the
TC parameters where saved with active temperature stabilization,
the QubeCL will
automatically start to stabilize the temperature.
If it manages to do so, L5 switches off
and L6 (green) switches ON.
• Once you connect the USB cable to the QubeCL, LED L1 (green) switches ON.
• Start the QubeCL_control
program,
wait until
it detects the connected QubeCL(s),
chose the one you want to connect to and press the blinking CONNECT button.
• The first thing to do when starting to work with a new set-up,
is to set all
the safety
parameters offered by the QubeCL:
– In the configuration panel,
set the maximum current limit following the max-
imum current limit indicated in your
laser
device datasheet and/or
test report.
Once set,
if the user tries to set a current higher than this maximum value,
the
software will
cap it to the maximum value.
– In the TC Advanced section of the Advanced tab(Figure 13), that can be found
in the lower panel, use the TEC limit control
to set the maximum current that
the TC controller may source to the Peltier stage.
To do so,
you must press the
correspondent CHANGE button.
Figure 13:
The QubeCL_Control
software Advanced tab
Once the safety values are set, the user may start to operate the QubeCL.
Before turning on the current driver, it is necessary to stabilize the temperature of the laser.
This can be done by the QubeCL if the TC module is included in the device configuration.
23 - 69

---
5.3.1
Temperature control
operation
This sequence of operation has to be performed only if the QubeCL houses a TC module,
being in charge of
temperature stabilization of
the laser.
All
the necessary controls can be
found in the Temperature Controller tab, in the lower panel
of the GUI.
• In the Temp SetPoint form,
set the temperature at which the laser device has to be
stabilized.
QubeCL accepts value with a precision down to the second decimal
place.
The temperature setpoint can be changed at any time during the operations.
• Set the I and D parameters of the PID controller to 0.0 and P to a middle value.
• Press the Tstab button to activate the temperature stabilization:
– Once done,
TStab button turns orange,
LED L5 lights red indicating that the
desired temperature has not been reached yet.
– At first, check the temperature plot on the GUI, if the measured temperature
drifts towards the desired setpoint, the sign of the loop is the correct one.
Other-
wise, if the measured temperature drifts away from the setpoint, it is necessary to
switch the sign of the loop using the dedicated sign buttons.
– Increase the I gain searching for an oscillating behavior of the measured temper-
ature around the setpoint.
Tstab button will become green each time the measured
temperature stays in a range of 20 mK around the setpoint.
Similarly, LED L5 will
turn off and LED L6 will
turn on in this range.
– Adjust P and I
parameters trying to reduce the oscillating behavior as much
as possible,
while also reducing the time needed to the measured temperature to
converge to the setpoint.
At last, TStab must be constantly green, LED L5 turned
off and LED L6 turned on.
Once the Temperature Controller parameters have been set to the optimal ones, it is possible
to store them in the QubeCL ROM memory using the Save button in the Temperature
Controller tab to have them as default from now on.
Remember that the save button will also
store the stabilization status,
meaning that if
the Temperature Controller were active while
saving, it will
automatically start to stabilize the laser once the QubeCL is switched on again.
To operate the Temperature Controller at best of its performance, please keep in mind the
following characteristics:
• If
the heatsink of
the laser
is unable to dissipate enough heat,
its temperature will
gradually raise.
If
so,
the Temperature Controller
will
have to correct an increasing
temperature difference,
being forced to constantly,
even if
slowly,
increase its current
output,
eventually reaching the maximum value.
Once reached,
the Temperature
Controller will
be unable to further stabilize the temperature of the laser.
To avoid this behavior, it is sufficient to check if the current sourced by the temperature
controller is not constantly drifting towards one of its maximum values.
Otherwise,
a
better heat management would be necessary.
24 - 69

---
• The Temperature Controller includes a safety behavior in order to avoid to operate
with a not stabilized laser for long time.
Once the temperature stabilization is activated,
the Controller constantly check the difference between the measured temperature and
the setpoint.
At each check, the difference is summed to the previoulsy measured ones.
If the measured temperature crosses the setpoint, the sum is set back to 0.
If the sum goes over a certain threshold,
the Temperature Controller raises an error,
meaning that is was not capable of stabilize the laser Temperature.
The threshold can
be chosen with the ErrorSumLim parameter which is present in the Advanced Tab.
This limit can be increased if it is too short and causes errors to be raised even when the
Controller is converging toward stabilization.
An higher threshold would nevertheless
increase the time before an error raises when the Temperature Controller is effectively
failing to stabilize the laser temperature.
5.3.2
Current control
operation
Once the laser temperature is stabilized,
the user can start to source current to it.
The
necessary controls are grouped together in the current driver panel
of the GUI.
• The desired current value can be typed into the form present in the Current Setpoint
section of the panel.
The value of the current setpoint can also be changed with the up
and down arrows or with the slider close to the form.
The buttons and the knob in the Fine Current Tuning tab enable the user to finely
control
the current setpoint value.
The Output Current setpoint may be changed at any moment during the operations.
• Once the setpoint is set, the current generator can be switched on by pressing the Iout
button.
A click noise must be heard coming from the QubeCL,
while the sourced
current plotted on the GUI
(purple line) gradually rises towards the setpoint (orange
line).
The QubeCL constantly monitors the value of the current sourced to the laser and the
value of the voltage across the laser.
The two values are shown in the Readings panel
along with the Power Source voltage monitor.
• If a modulation of the laser current is needed, it can be applyed by enabling the ModEn
button, which in turn will give the user the ability to switch on the M1 and M2 buttons.
In order for the modulation to be sourced to the laser, there must be a modulating signal
connected to the dedicated inputs of the QubeCL or a signal
generated by an onboard
DDS module.
This matter is examined in detail
in a subsequent chapter.
The ModEn button is not active for 10 seconds starting from the current being switched
on, so that the modulations can be applied only when the current to the laser is stable.
Modulations will
be automatically switched off when the current is.
The parameters for the Current Generator can be stored in the ROM memory of the QubeCL
to be used as default when the QubeCL is switched on by clicking the Save button.
Unlike the
Temperature Controller, save button of the current control
panel
will
not save the Current
Generator status,
hence the QubeCL never automatically starts to source current to the
laser when it gets switched on.
25 - 69

---
5.3.3
Automatic protections
QubeCL continuously monitors a variety of functional
parameters both of the laser and of
the QubeCL driver
itself.
The data acquired allows the QubeCL to implement a series of
automatic protections to operate as safe as possible,
avoiding damage to the QubeCL and,
most importantly, to the laser device.
The QubeCL driver implement the following protection for the Current Generator:
Open-circuit protection:
this safety measure con-
stantly checks if the electrical
connection with the
laser is still
present,
otherwise the Current Gener-
ator
is immediately switched off and the QubeCL
output is shorted.
If this condition is detected, a blinking warning sig-
nal
appears on the GUI,
next to the Vlas indicator
in the Readings box of the Current Generator panel.
Compliance Limit reached:
the same warning described above is raised if
the QubeCL
senses a voltage across the laser higher than the maximum allowed compliance voltage.
Such
Compliance Limit [CL] voltage depends on the nominal
Source Voltage for the Current Gen-
erator and can be obtained as CL = (0.77 ·Vcc) - 1.
Low Voltage Supply protection:
if the CM mod-
ule supply voltage falls below 18.5 V,
the QubeCL
automatically switches off the current and shorts the
output connector.
A blinking warning signal
appears on the GUI, next
to the Vcc monitor in the Readings panel.
The previously described errors will
not be deasserted if
the warning condition is not met
anymore (for example,
if
the Supply Voltage raises again above 18.5 V).
This behavior has
been implemented to avoid the warning to go unnoticed in case the user were not monitoring
the GUI or the QubeCL_Control program were not running when the error condition happened.
Once the cause of the error condition has been eliminated,
the user can reset the errors by
switching on the current generator.
26 - 69

---
The QubeCL also monitors the temperature of the laser,
implementing some safety features
to prevent potentially dangerous conditions:
TEC module communication fault:
if some problem oc-
curs to the communication with the TEC module (i.e.
when
the module is not properly powered) the Temp Read value
becomes -273 °C, and a warning box appears next to it (see
Figure at right).
The warning disappears as soon as the com-
munication with the TEC is re-established.
Temperature sensor fault:
in case the temperature sensor
is not working properly (i.e.
it is not connected to the TEC
module), the Temp Read value becomes -67.737 °C and the
warning square box appears next to the readed value.
In case one of the above problems occur when the Temper-
ature Stabilization is active, a warning box appears also next
to the Temp Read line.
If
the QubeCL is sourcing current
to the laser when a Temperature Controller fault is detected,
the current source is disabled to prevent damage to the laser.
Temperature stabilization problem:
if,
for
any reason,
the Temperature Controller modules fails to stabilize the laser
temperature (i.e.
wrong PID sign,
poor heat management),
the QubeCL disables both Temperature Stabilization and Cur-
rent output to the laser.
A warning box appears next to the
Temp Read line.
The threshold above which a Temperature stabilization is
raised can be changed by acting the Error
Sum Lim pa-
rameter in the Advanced tab, as previously described.
27 - 69

---
6
Advanced Operations - Modulators and DDS
Besides the basic functionality,
QubeCL may perform a variety of advanced operations while
driving a laser device.
As well
as the basic functions,
the advanced ones are performed by
dedicated modules.
The laser Head (LH) module is the one deputed to the connection to an external
laser and,
besides providing all
the necessary switches and safety features,
it also includes two current
modulator circuits.
The modulator circuits onboard of the LH are capable of directly injecting current to the laser,
superimposing their output signal over the DC bias current sourced by the Current Generator.
The modulating current signal generated by the modulators is proportional to the input signal
that is applied to the dedicated inputs on the LH module.
If we denote the modulation current
with I
,
the input control
voltage with V
and the absolute value of
the gain in mA/V
mod
in
with G, the modulation current is equal
to:
I
= −G · V
.
mod
in
This means that
a positive control
voltage will
subtract current from the laser
bias
current while a negative control
voltage will
add the modulating current to the DC bias.
The two modulators can be activated ad the same time,
in which case their output signals
would arithmetically add together.
The two modulators are built in order to cover all
the possible needs that may arise while
driving a laser:
• Low current modulator:
this modulator has a lower V/I
Gain.
It allows the user to
modulate the bias current with signals down to tens of nA with great precision and low
added noise.
• High current modulator:
this modulator has higher Gain,
allowing to modulate the
Bias current with signals up to tens of mA, allowing the user to perform large scans.
The gain values for the two modulators can be found in the QubeCL datasheet and in the test
report that comes along with the instrument.
CAUTION
The high current modulator can provide currents up to tens of mA which can be added
to or subtracted from the bias current of the laser.
When sourcing negative control
voltages to the modulators,
the modulating current
is added to the bias current,
in this case the total
current may exceed the maximum
current allowed by the laser.
Therefore it is necessary to be very careful
to avoid the
risk of damaging the laser itself.
The maximum current value that the user can set on the QubeCL does not comprehend
the modulating signal,
being just referred to the sourced Bias Current.
The QubeCL
does not perform any safety check on the overall
Output Current value resulting as the
Bias Current with the superimposed modulation signals.
28 - 69

---
Figure 14:
Appearance of the LH module equipped with the DDS sub-module
6.1
The DDS sub-module
The LH module may host a smaller sub-module, namely the DDS sub-module.
DDS stands
for Direct Digital Synthesis, being a digitally controlled electronic circuits capable of generating
analog signals
with the ability of
controlling some of
their
parameters
(being Amplitude,
Frequency, Phase and Waveform).
This sub-module adds two DDS circuits to the LH module, enabling the QubeCL to generate
its own modulating signals without needing an external
signal
generator
connected to the
Modulation inputs.
Since the DDS sub-module is housed inside the LH module,
it won’t
modify the physical
dimensions of the QubeCL driver.
The generated signals are directly injected into the DC bias current similarly to the signals
from the externally sourced modulators.
When the QubeCL also includes a PDH or
LIA module,
the DDS sub-module is capable
of
interacting with them,
providing all
the synchronous signals that are necessary for those
two modules to perform their operations.
For more information,
please check the chapters
dedicated to the PDH and LIA modules.
6.1.1
Connectors
The presence of
the DDS sub-module adds some feature to the QubeCL and consequently
some minor modification to the LH module interconnections are present.
The three standard connectors previously described are still
present, but the connector C7 is
added on the left side of the module, as it is shown in Figure 14.
C7 is an output SMA connector.
The DDS module is capable of generating a 5 V square-
wave synchronization signal
which is sourced from the low impedance output on C7.
The
user may chose to synchronize this signal
both to the waveform sourced by DDS1 or to the
waveform sourced by DDS2.
6.1.2
Control
Software interface
The QubeCL _Control Software includes a section dedicated to the control of the DDS module
functionalities.
The section can be accessed in the lower part of the GUI by selecting the DDS
tab, which is shown in Figure 15.
The DDS control
tab is subdivided into three panels:
• The Sync.
Channel panel includes the controls for the external synchronization signal.
• The Channel 1 and Channel
2 panels include all the interfaces to properly control the
characteristics of the signal
generated by the two DDS circuits.
29 - 69

---
Figure 15:
The DDS control
panel
of the QubeCL_Control
Software
The Sync.
Channel panel allows the user to select to which DDS the external synchroniza-
tion channel
must be synchronized.
To do so, the GUI shows the Sync.
frequency selector.
Below the selector, a read-only form shows the actual frequency of the synchronization signal,
which is equal
to the one of the selected DDS channel.
Above the Sync.
Channel
panel,
the Save param.
button allows the user to store the
current parameters into the ROM memory of the QubeCL so that they can be restored at the
next switch on of the driver.
The Channel
1 and Channel
2 panels contain the same controls,
replicated for the two
DDS circuits:
• ENABLE: this button activates the DDS circuit, sourcing the synthesized signal
to the
laser.
Note that the DDS circuits share the same signal paths of the external modulators,
hence it is necessary to activate the modulators from the Current Control
panel
in
order for the DDS signal
to be sourced to the laser.
• WAVEFORM: this control
allows the user to chose the waveform of the signal
synthe-
sized by the DDS circuit.
The available options are Sine wave and Triangular wave.
• AMPLITUDE:
this form allows the user
to set
the peak-to-peak amplitude of
the
modulating signal,
expressed in mA.
The value can also be modified by using the up
and down arrows next to the form.
• FREQUENCY: allows the user to set the frequency of the modulating signal.
• PHASE: allows the user to set the phase of the modulating signal.
30 - 69

---
6.1.3
DDS operations
To operate with the DDS module, follow the steps listed below:
• Activate the Temperature Controller (if present) and the Current Generator as previously
described in this manual.
Wait until it is possible to interact with the Modulation Enable
controls on the Current Control
panel.
• Activate the overall Modulation Enable button and the enable(s) for the desired mod-
ulation channel(s).
• In the DDS tab, set the desired parameters for the modulating signal
(Amplitude, Fre-
quency, Phase, Waveform) and activate the DDS with the ENABLE button.
Even when the DDS sub-module is present, it is still
possible to modulate the laser cur-
rent with the external modulators.
The internally generated signal(s) and the externally
applied one(s) will
be added together and injected over the DC bias current.
• The parameters of the synthesized signal
can be changed at any time during the oper-
ation.
• If needed,
the External
Synchronization Signal
can be used as a trigger source for
an oscilloscope, to synchronize its acquisition with one of the DDS generated signals.
• Disable the DDS-generated signal
by deactivating the corresponding ENABLE button
when it’s not needed anymore.
Detailed instructions on how to use the DDS sub-module in pair with the PDH or LIA modules,
if present, can be found in the respective chapters.
31 - 69

---
7
Advanced Operations - PLL module
The PLL module gives to the QubeCL instrument the capability of easily perform a complex
task such as the phase-lock of the controlled laser frequency to a reference source.
The PLL module requires two input signals:
the radio-frequency (RF ) signal
generated by
the beatnote between the controlled laser and the reference laser and the signal from a Local
Oscillator (LO) used as a set-point for the RF signal.
When the PLL is activated, a current proportional
to the phase error signal
is processed by a
Proportional-Integral
(PI ) stage and internally added to the bias current in order to close
the phase-stabilization loop.
The RF and LO frequencies can be divided by an integer value by the mean of a frequency
divider,
in order
to obtain two signal
with similar
frequencies at the input of
the internal
phase-detector.
Each divider can be independently set using the Control
Software,
while the
two divided signals can be monitored on the DIVIDER MON output (C10 ).
When the loop
is activated,
the PHASE mon output (C11 ) provides a real-time replica of the phase-error
signal
after the PI stage.
Since each laser can only be locked to a reference at a time,
in this case another laser
wavelength, the PLL Module can not be used with other Locking Modules (LIA, PDH) on the
same QubeCL system.
7.1
Connectors
Figure 16:
Appearance of the PLL module
C8.
RF in
Signal
level :
C10.
DIVIDER
Requires 50 Ω load
intput
-30 dBm 0 dBm
mon
AC coupled.
bandwidth:
output
10 MHz
300 MHz
C9.
LO in
Signal
level :
C11.
PHASE mon
Requires Hihg-Z load.
intput
-30 dBm 0 dBm
output
DC coupled.
bandwidth:
Range:
10 MHz
300 MHz
-1 V / 1 V
7.2
Control
Software Interface
The QubeCL Control
Software has a dedicated section for the control
and monitor of the
PLL Module, it can be accessed by clicking on the PLL tab in the lower portion of the GUI.
32 - 69

---
As can be seen in figure 17,
the panel
dedicated to the PLL module is divided into two
main areas:
on the rightmost part there are two panels named Phase Error Plot and Transfer
Function which in turn show the actual
error signal
monitored by the PLL module and the
transfer function implemented by the PI
loop mounted on the module itself with the current
parameters.
The remaining part of the panel
is dedicated to all
the controls available for the
user to control
the behavior of the PLL Module.
Figure 17:
Control
Software Interface dedicated to the PLL Module
The controls for the PLL module are subdivided into categories based on their functionality.
The Phase Detector
Settings panel
gives the user
access to the controls for the Phase
Detector:
• CPG: sets the output gain for the Phase Detector, has a range between 1 and 8.
It has
a proportional
effect on the transfer function of the module.
• GPG mult:
works as a multiplier
for
the CPG control,
the total
proportional
gain
introduced is the product of the two.
• LO Divider:
sets the frequency division factor (integer) for the Local
Oscillator signal.
• RF Divider:
sets the frequency division factor (integer) for the RF input signal
(e.g.
the beatnote).
• Sign:
sets the sign of the PI correction loop.
The Divider Monitor panel
gives the user the ability to chose what signal
will
be sourced
to the Division Monitor Output connector (C10).
When in OFF position,
no signal
will
be
available.
The other two buttons allow to chose between the post-divisor Local
Oscillator
signal
or the post-divisor RF signal, so that the user can visualize the actual
frequency of the
two signal
once they’ve been divided.
33 - 69

---
The Lock panel
allows the user to chose the functional
mode of the correction PI
loop of
the PLL module.
There are three possible options:
• OFF:
the PI
loop is not working and does not add any correction signal
to the laser
current.
• P:
the loop works in Proportional
mode,
only the proportional
gains affect the output
correction signal.
• PI:
the loop work with both Proportional
and Integrative components,
being fully en-
abled.
On the central
bottom part of the tab, the Phase PI settings allow the user to control
the
PI loop gains:
• Phase Prop:
sets the Proportional
Gain of the PI loop.
• Phase Int:
sets both the Integrative Gain of the PI loop and the high frequency cut-off.
The Transfer Function plot gets automatically updated each time that any of the previously
mentioned control is changed, with the exception of the Division Monitor which does not affect
the transfer function of
the PI
loop.
For a correct visualization of
the loop,
the user must
input the laser Frequency tuning into the dedicated form on the plot panel.
The Phase Error Plot shows the real
time value of the phase error.
The sampling rate is
low, hence it will
not be helpful
with fast moving signals, but can give the user and easy way
to monitor the time evolution of this signal
during the locking.
7.3
PLL Operation
For a proper set-up and use of the PLL module, please follow the steps listed below:
• Check that the Beatnote (RF ) signal
has a level
between -30 dBm and 0 dBm and a
frequency between 0 and 300 MHz.
This signal
must be connected to the RF IN (C8 )
conenctor.
• Use a function generator to generate a Local
Oscillator
signal
with a level
close to 0
dBm and a frequency between 10 MHz and 100 MHz.
This signal
must be connected
to the LO IN (C9 ) connector.
• Set both LO divider and RF divider to 1.
• Connect the DIVIDER MON output connector (C10 ) to an oscilloscope with a 50 Ω
termination.
By using the Div Mon control, switch the signals that have to be sourced
on this port.
The output signal
must be a clean square wave with the same frequency
of the selected input signal
(RF or LO).
If the output frequency is not properly counted, a better filtering is required for the input
signal, in order to obtain an higher SN ratio for the counter to properly work.
34 - 69

---
• Set the values of
LO divider and RF divider so that the two subdivided frequencies
are as close as possible.
For example:
if f
= 100 MHz and f
= 10 MHz, set LO div = 1 and RF div = 10.
RF
LO
• Write the proper value for the laser Tuning Coefficient in the Laser Tuning input form
in the Transfer Function Plot tab.
Then set the values for CPG, CPG mult, Phase
Prop and Phase Int aiming to obtain a Transfer Function Plot that properly crosses
the Unity Gain level
to obtain stability.
• Check the Phase Error Plot to see if the Phase Detector is properly working:
finely
tune the bias current from the Current Control
Panel
in order to force a change in the
beatnote frequency.
Make the RF signal frequency cross the LO one to see if the Phase
Error switches from positive to negative or vice versa.
• Tune the Bias Current in order to approach the desired frequency for the lock and enable
the Proportional
Lock loop (P button).
If the RF signal
frequency sweeps away from
the LO one, change the Loop Sign, otherwise the PLL is already working to lock the
signal.
• Enable the PI button of the Lock panel.
• Adjust the CPG,
Phase Prop and Phase Int settings in order to optimize the lock
quality.
More than one locking condition can be found,
with different bandwidths and
gains.
Once the lock has been successfully achieved, the Slow Loop functionality can be activated
in order to compensate phase errors arising because of slow drifting environmental parameters.
To do so, please check chapter 10.
35 - 69

---
8
Advanced Operations - PDH module
The PDH module gives to the QubeCL instrument the capability of easily perform a com-
plex task such as the lock of the controlled laser frequency to an optical
cavity.
The PDH module requires two input signals:
the radio-frequency (SIGNAL IN) signal
gen-
erated by the frequency-modulated laser beam reflected back from the cavity and the Local
Oscillator (REF IN) signal
used for demodulating RF signal.
The LO is usually generated by
the same function generator used for the frequency modulation, after a careful
tuning of the
relative phase.
When the PDH is activated,
a current
proportional
to the error
signal
is processed by a
Proportional-Integrative (PI ) stage and internally added to the bias current closing the frequency-
stabilization loop.
Since each laser can only be locked to a reference at a time, in this case an optical cavity, the
PDH Module can not be used with other Locking Modules (LIA, PLL) on the same QubeCL
system.
8.1
Connectors
Figure 18:
Appearance of the PDH module
C12.
Signal
in
Signal
level :
C14.
MON OUT
Requires Hi-Z Ω load
input
-30 dBm 20 dBm
output
DC coupled.
bandwidth:
Range:
300 kHz
100 MHz
-1 V / 1 V
C13.
REF in
Signal
level :
C15.
HOLD
Signal
level:
input
7 dBm
input
TTL.
bandwidth:
Bandwidth:
300 kHz
100 MHz
10 MHz
8.2
Control
Software Interface
The QubeCL Control Software has a dedicated interface for the control of the PDH module,
which can be accessed by clicking on the PDH tab in the lower half of the GUI.
As it can be seen in Figure 19,
the PDH control
tab is subdivided into two main parts,
as
well
as the PLL tab is.
On the right side, two monitor plots can be selected in the respective
tabs, so that the user can have a look on the transfer function of the PI
loop of the module
or at the level
of the correction signal
produced by the PI loop itself.
On the left side, the tab includes all
the necessary controls to properly use the PDH Module.
36 - 69

---
Figure 19:
Control
Software Interface dedicated to the PDH Module
On the top right corner of
the PDH tab,
there is the PDH Lock set of
controls,
which
allows the user to perform various actions:
• The first couple of buttons, P and PI, allow the user to chose the mode of the correc-
tion loop of the PDH module,
switching between pure Proportional
and Proportional-
Integrative.
• The LOCK button allows the user
to activate the correction loop by injecting the
correcting current on the laser.
• The Hold button enables the HOLD input (C15 connector).
When the Hold Enable button is active, the automatic correction loop of the PDH module
can be put into an Hold State in which is it closed on its set-point, so that its output remains
constant at the level
it was when the Hold signal
is asserted.
To externally control
the Hold
function, a TTL signal
must be connected to the C15 connector.
A 5 V level
would activate
the correction loop while a 0 V level
leaves it in an Hold state.
Next to the PDH Lock panel, the Save button allows the user to store the current settings
for the PDH module into the ROM memory of
the QubeCL,
so that they will
be used as
default the next time the driver is switched on.
The PDH Mon panel
contain a switch button that allows the user to chose which signal
will
be present on the Monitor Out (C14 connector).
The available options are ER which
will
show the Error signal
received by the module from the PDH detector, and LM which will
show the correction signal
generated by the correction loop on board of the PDH module.
Below the PDG Mon,
there is a panel
containing all
the controls for the regulation of the
Input Offset for the Correction Loop of the PDH Module.
The desired offset can be directly
typed into the dedicated form,
otherwise the four up and down arrows can be used.
Each
arrow will
add or remove an offset value correspondent to the value written in the form below
the couples of arrows.
37 - 69

---
The PDH PID settings panel
allows the user to tune the parameters of the PI Correction
Loop.
• PRE AMP allows the user to set the gain of the pre-amplifier of the loop.
It accepts
value between 2 an 11.
This value is a pure gain, hence it has no unit of measurement.
• PDH P1 allows the user to set the overall
proportional
gain of the loop.
• PDH P2 allows the user to set the proportional
gain of the loop,
but also influences
the position of the integratife cut-off frequency when the loop is in PI mode.
• PDH I
allows the user
to set the cut off frequency of
the integrative action of
the
correction loop.
The Slow Loop control
panel
is identical
to the one that is present in the PLL tab, users
can refer to the dedicated chapter 10.
The Transfer Function Plot gives a visual
representation of the transfer function of the
correction loop.
The plot automatically updates each time the user change one of the param-
eters for the loop itself.
For a correct visualization of the plot, it is necessary to fill
the Laser
Tuning form with the actual
parameter of the laser being driven by the QubeCL.
The PDH Lock Mon Plot shows the temporal evolution of the correction signal generated
by the PI
loop of the PDH module.
The digital
monitor used to update this plot has a low
sampling frequency, hence it can’t be used to monitor high frequency signals on screen, but it
gives a long history of the locking signal
evolution during time.
8.3
PDH operations
For a proper set-up and operation of the PDH module, follow the steps listed below:
• Verify that the signal
reflected from the cavity has a level
between -40 dBm and 20
dBm, with a frequency between 300 kHz and 100 MHz.
Connect this signal
to the SIGNAL in (C12 ) connector.
• Use a function generator to generate a proper demodulation signal, with 7 dBm magni-
tude and a frequency between 300 kHz and 100 MHz.
Connect this signal
to the REF in (C13 ) connector.
• Activate the modulation signal
to the laser,
monitoring the ER signal
youll
be able to
find the characteristic PDH signal
waveform.
Adjust the modulation signal
phase until
you find the value that minimizes the carrier contribution to the signal, leaving only the
sidebands contributions.
At that point,
the right phase to obtain a good lock is that
value ±90°.
Try one of
those two values and LOCK the loop in P mode,
if
only the
sidebands contributions are magnified try the other phase value,
otherwise this is the
right one.
38 - 69

---
• While locked and in P mode, change the P1 and P2 values until you find the combination
that produces the best effects on the signal
transmitted by the cavity.
Reduce the
amplitude of
the ramp signal
used to scan the lasers frequency searching for the one
that matches the cavity while adjusting its offset to keep the lasers frequency close to
the resonance.
• Activate the PI
mode and search for the best PI
parameters combination,
which guar-
antees the best locking performance.
Once the lock has been successfully achieved, the Slow Loop functionality can be activated
in order to compensate phase errors arising because of slow drifting environmental parameters.
To do so, please check chapter 10.
39 - 69

---
9
Advanced Operations - LIA module
The LIA (Lock-In Amplifier ) module gives the QubeCL the ability to perform lock of the
controlled laser frequency to the first or second derivative of a molecular absorption line profile.
The LIA requires two signals to operate:
one demodulating signal
and the signal
from the
detector used to retrieve the absorption signal
from the reference molecular sample.
The de-
modulating signal must be synchronous with the modulating one.
To obtain such synchronous
signals, the LIA module must work in pair with the DDS sub-module, the latter being
in charge of generating both the modulating and demodulating signals.
The modulation-demodulation action performed by a QubeCL equipped with a LIA module
allows the instrument to obtain a signal
proportional
to the first or second derivative of the
absorption line signal retrieved by the detector.
This signal can be used as an error signal to be
processed by an internal Proportional-Integral (PI) Correction Loop.
Doing so, LIA module can
generate a correction current which is directly fed to the laser,
closing the frequency-locking
loop over the reference absorption line.
Since each laser
can only be locked to a reference at
a time,
in this case a molecular
absorption line, the LIA Module can not be used with other Locking Modules (PLL, PDH) on
the same QubeCL system.
9.1
Connectors
Figure 20:
Appearance of the LIA module
C16.
Signal
Out
Requires Hi-Z load.
C18.
Sine Out
Requires Hi-Z load.
output
DC coupled.
output
DC coupled.
Signal
level :
Signal
level :
-1 V 1 V
-1 V 1 V
C17.
Signal
In
AC coupled.
C19.
Ramp Out
Requires Hi-Z load.
input
Signal
level:
output
DC coupled.
< 1 Vpp.
Signal
level :
-1 V 1 V
9.2
The DDS sub-module operation in pair with the LIA module
The operation of
the LIA modules require the modulating and demodulating signals to be
synchronized one to the other.
To achieve perfect synchronization, the use of a single timing
40 - 69

---
source is required.
The DDS sub-module provides such timing source,
generating both the
modulating signal
with the DDS1 and the demodulating signal
which is internally fed to the
LIA module
The demodulating signal
frequency is fixed at 32.768 kHz,
hence the modulating signal
frequency is forced to operate in two possible states:
• To retrieve the first derivative of the molecular absorption line,
the modulating signal
must have the same frequency as the demodulating one, hence the DDS1 frequency is
fixed at 32.768 kHz.
• To retrieve the second derivative of the molecular absorption line, the modulating signal
must have half the frequency of the demodulating signal, being 16.384 kHz.
In both cases,
the DDS1 frequency cannot be changed by the available control
in the DDS
panel, which becomes inactive.
The reciprocal
phase between the modulating and demodulating signal
is another crucial
factor to achieve a good lock with the LIA module.
Because of
that,
if
the LIA module is
set to be active at the startup of the instrument, it performs a rephasing process in order to
obtain a known initial
status of
the phase between the two signals.
The process runs for a
few seconds until
the phase between the two is equal
to the DDS1 phase saved in the ROM
memory of the QubeCL.
The LIA module is ready to be used once the rephasing process is
done.
If
the LIA module is not necessary and the user wants to have full
control
on the DDS1
frequency, the LIA demodulating mode can be set to OFF (see the below chapter), to remove
the frequency constraints for DDS1.
The rephasing process too is not performed at startup if
the LIA demodulating mode is set to OFF.
41 - 69

---
9.3
Control
Software Interface
Figure 21:
Control
Software Interface dedicated to the LIA Module
The QubeCL Control
Software section dedicated to the LIA module can be accessed by
clicking on the LIA tab in the lower part of the GUI. Similarly to the PLL and PDH modules,
also this control
tab is subdivided into a monitor section and a control
section.
On the center portion of
the LIA tab,
the user can find the commands that control
the
functionality of the correction loop included into the LIA module:
• The Loop mode switch allows the user to chose if the Control
Loop should work in
pure proportional
mode or in proportional-integrative mode.
• The Filter enable button enables or disables a tunable Low-Pass filter that may be
useful
to smooth out the Correction Signal
generated by the Loop.
• The LOCK buttons enable or disables the correction signal
fed to the laser, effectively
closing the correction loop.
Next to this section, user can find the Save button.
By pressing it, the user will save all the
current LIA parameters an functional
modes into the ROM memory of the QubeCL so that
they’ll
be used as default the next time the QubeCL is switched on.
The ON button to the left of the save button allows the user to switch off the LIA operations,
deactivating the demodulator.
The Phase Ready LED indicates if the rephasing process between the DDS and the de-
modulating signal
is completed.
42 - 69

---
The LIA Loop Settings box holds all
the controls to tune the PI Correction Loop gains:
• VGA GAIN set the Gain of
the Variable Gain Amplifier (VGA) being the first stage
of the Correction Loop.
The gain is expressed in dB.
The value of the Gain can be
changed both by direct typing or with the up and down arrows.
Accepted values range
from 0 dB to 80 dB.
• Prop.
sets the magnitude of
the Proportional
action of
the Correction Loop.
This
control also affects the cut-off frequency of the Integrative component of the Correction
Loop.
• Integ.
sets the cut-off frequency of the Integrative component of the Correction Loop.
• Filter sets the -3 dB frequency for the Low Pass Filter that can be activated with the
Filter Enable button.
The leftmost section of the LIA tab holds all
the control
specific for the Lock-In Amplifier
and the monitor selector:
• The Monitor Selector switch allows the user to decide which signal
is present on the
SIGNAL OUT (C16 ) port of the LIA module.
By selecting LOCK,
the LIA module
will
output the correction signal
generated by the internal
PI
control
loop.
The ERR
selection sets the error signal
(which is the input to the PI
control
loop) to be on the
SIGNAL OUT connector.
• The Digital Filter drop-down menu allows the user to select one of the available built-in
digital
filters.
More information about the specific behavior of each filter can be found
in the dedicated chapter below.
• The Demod.
Freq.
switch allows the user to select in which mode the LIA module
should works.
When f is selected, the LIA works to lock the laser on the first derivative
of
the absorption line,
while using the 2f
mode sets the LIA to work on the second
derivative of the absorption line.
When the OFF button is selected,
the LIA is set to
not work in pair with the DDS module.
9.4
LIA operations
For a proper set-up and operation of the LIA module, follow the steps listed below:
• If necessary,
set the desired demodulating mode and wait up to a few seconds for the
rephasing process to be completed.
The Phase Ready LED will
light up once the
process is completed.
• Connect the detector to an oscilloscope and use the DDS2 to scan the laser wavelength
to find the absorption line of the reference molecule with a low frequency triangular wave
modulation.
To easily achieve this condition, lock the external synchronization signal of
the LH module to the DDS2 and use it as the trigger for the oscilloscope acquisition.
43 - 69

---
• Once the absorption line has been found and it is well
centered in the scope screen,
activate the DDS1 and set a proper modulating amplitude.
The LIA module is capable
of
demodulating signal
down to a few tens of
mV,
higher signals risk to saturate the
internal amplifier stages.
Check the signal from the detector to verify that the amplitude
of the modulation is in this accepted range.
Connect the signal
from the detector to
the SIGNAL IN connector of the LIA module.
• Connect the SIGNAL OUT connector to the oscilloscope in order to have a monitor
signal
to be used as a reference during the locking operation.
The oscilloscope must
have a high impedance input.
Set the monitor selector to ERR. The first derivative of
the absorption line must appear on the scope,
as a consequence of the demodulating
action performed by the LIA.
• Set the Digital Filter that produces the best signal quality and maximizes the amplitude
of the signal
on the scope.
• By adjusting the modulating signal
phase, search for the worst condition (lowest signal
on the monitor).
Once achieved, change the modulating frequency phase by 90°to obtain
the optimum phase condition.
It may be necessary to switch between the +90°and -
90°phase conditions in order to find the proper correction loop sign.
• Set the correction loop to P mode and close the correction loop with the LOCK
button.
Activate and tune the Low Pass Filter if necessary.
• Adjust the Prop.
and VGA gains to maximize the correction signal
without occurring
in an oscillating behavior.
• Set the correction loop to PI mode and adjust the Integ.
time constant to achieve a
good locking.
Once the lock has been successfully achieved, the Slow Loop functionality can be activated
in order to compensate phase errors arising because of slow drifting environmental parameters.
To do so, please check chapter 10.
9.5
LIA digital
filter
The LIA module is equipped with a mixed signal
Lock-In Amplifier which includes a set of
built-in digital
filters.
The digital
filters shows a frequency behavior which is dependent on
the clock frequency used to drive the filter.
In the LIA module, the clock signal
used for the
digital
filter is generated by the DDS sub-module with a fixed frequency of:
20
f
= 2
= 1.048576MHz
CLK
The clock signal passes through a series of frequency divider before reaching the programmable
digital
filter, thus obtaining a frequency of:
17
f
= 2
= 131.072kHz
SO
The available filters,
which can be selected by the mean of the dedicated drop-down menu,
are the following ones:
44 - 69

---
• BP0:
Band Pass filter, centered on f
/8 = 16.384 kHz
SO
• BP1:
Band Pass filter, centered on f
/4 = 32.768 kHz, Q = 8.4
SO
• BP2:
Band Pass filter, centered on f
/4 = 32.768 kHz, Q = 4.3
SO
th
• LP1:
Low Pass filter, -3 dB at f
/5 = 26.2144 kHz, 4
order
SO
th
• LP2:
Low Pass filter, -3 dB at f
/8 = 16.384 kHz, 4
order
SO
st
• Notch:
centered on f
/4 = 32.768 kHz, 1
order
SO
• All
Pass
45 - 69

---
10
Advanced Operations - Slow drift compensation
All
three of
the locking modules previously described are built to com-
pensate fast drifts in the laser wavelength in order to lock it to a specific
reference, depending on the locking mechanism which is being deployed.
Nevertheless,
a variety of environmental
factors (such as temperature varia-
tions,
mechanical
vibrations and strains,
optical
feedbacks,
etc.)
may intro-
duce slow drifts in the locking point.
To maintain the lock to the reference,
the PI
correction loop of the locking modules is forced to follow those slow
drifts in order to compensate them.
If the control
loop output drifts towards
one of its saturation levels,
the lock will
be lost due to the inability of the
PI
loop to further
correct the laser
wavelength against the environmental
disturbance.
To avoid this kind of
slow drifting interference to break the lock,
each
locking module may automatically compensate those slow drifts by directly acting on the
bias current fed to the laser
or
on the temperature setpoint of
the laser
being controlled
by the onboard TC module.
This correction loop can be activated and controlled with the
dedicated Slow Loop box which is present in each one of
the locking module tabs of
the
QubeCL_Control
software.
The Slow Loop Control
box includes the followings commands:
• Slow Loop button:
enables and disables the Slow Loop automatic compensation.
• Actuator:
allows the user to decide whether the slow loop acts on the Bias Current or
on the Temperature Setpoint of the laser.
• Sign:
sets the sign of the slow loop.
• SL cycle time:
sets the sampling rate (expressed in ms) for the slow loop to act.
• Iset Limit:
sets the maximum variation of the bias current sourced to the laser that
the slow loop can produce.
Once the lock has been successfully achieved (with whichever of the locking module is being
deployed on the instrument), follow these steps to activate the slow loop:
• Chose the operating mode of the Loop:
Current or Temperature.
Note that acting
on Temperature forces the Loop to have a very long response time from the laser, with
the risk of triggering an oscillating behavior.
• If
the Slow Loop is acting on the current of
the laser,
set the maximum allowed
correction current.
• Enable the automatic correction by pressing the Slow Loop button.
46 - 69

---
• Check the error monitor plot to see if the Slow Loop is working properly:
the error
must be drifting toward 0, otherwise the Sign of the Slow Loop must be reversed.
• Adjust the SL cycle time to make the compensation loop faster or slower, if needed.
47 - 69

---
11
Advanced Operations - QubeCL advanced settings
Figure 22:
The first Advanced tab of the Control
Software GUI
The QubeCL_control
Software also has 2 tabs for advanced settings.
The two tabs hold
a series of
controls that are less frequently useful
or which give a direct control
over some
specific functionalities of the QubeCL. The two advanced tabs are available in the lower half
of the Control
Software GUI.
11.1
Adv.
1 tab
The first advanced tab, visible in figure 22, is subdivided in five different panels, each one
dedicated to a particular set of commands and controls.
The Direct Communication panel regroups all the commands that can be used to directly
communicate with the QubeCL by textual
commands instead of the interaction with the But-
tons and Forms contained in the GUI.
The command that the user wants to send to the QubeCL must be typed into the COM-
MAND form,
a complete list of the available commands with their syntax can be found in
chapter 13.
Once the command has been sent by pressing the SEND COMMAND button, the reply (if
any) from the QubeCL can be read in the READBACK form.
The CM Advanced panel
regroups some information about the CM module(s) on board
of the QubeCL, which are the one(s) deputed to the generation of the Current for the laser.
Here, the temperature of the modules is shown in the CM Temp box, while DtLoop shows the
update time of the GUI. The update can be forced by pressing the Update Panel
button.
The Power
button allows the user to switch off the power supply for all
the modules except
the TC,
which has a separate power line.
Using this command to switch off some of
the
module while the QubeCL is operating is an advanced operation and may lead to unexpected
behavior, do not use it when the laser is connected to the QubeCL.
48 - 69

---
The TC Advanced panel
holds advanced commands for the TC module.
Similarly to the
CM Advanced panel,
there is a Power
button that allows the user to switch on and off the
power to the TC module.
As well
as for the CM module, this operation is dangerous and
must be handled with care.
TEC Mode switch allows the user to decide whether the TC module should act as a Tem-
perature Controller for the laser (strongly recommended option) or as an Heater.
TEC limit is the maximum current that the TC module can source,
in both directions,
to
the Peltier module mounted on the laser.
User can change this value by the mean of
the
CHANGE button,
which will
open a dedicated window.
Set the Maximum current accord-
ingly to the one stated in the datasheet or test report of the laser that needs to be driven by
the QubeCL.
The two TSetMAX and TSetMIN values are used to set the maximum and minimum tem-
perature that can be selected as setpoint for the Temperature Stabilization Loop operated by
the TC module.
As well
as for the TEC current limit, those values can be changed using the
respective CHANGE buttons.
The ErrSumLim parameter is the one used as a safety measure to prevent the Temperature
Controller to drift away from the setpoint or to fail
to stabilize the temperature for any other
reason.
A detailed explanation of this parameter can be found in chapter 5.3.3.
The System Modules panel
shows the composition of
the QubeCL by highlighting the
LED corresponding to each module present in the QubeCL. If there are more than one module
of each kind (CMs, for example), the number of modules is also plotted in the corresponding
LED.
The configuration can be changed by clicking on the CONFIG button,
however ppqSense
strongly discourage the modification of
the QubeCL stack by the customers to avoid any
damage and the warranty nullification.
The System Advanced panel
holds some rarely used commands.
The Fan Control
com-
mand allows the user to chose the speed of the cooling fan,
ppqSense strongly recommends
to leave this selector on the Auto position.
The two buttons on the lower part of the panel
can be used to activate the Wi-fi or the LAN
connection, if the QubeCL is provided with such options.
The two forms in the lower right part of the panel
can be used to chose how often the data
acquired from the Control
Software have to be saved in the log file.
Lastly, the ProgOn button which allows to use the Control
Software to update the firmware
of the QubeCL. This operation is dangerous and must never be executed when the QubeCL is
connected to a laser.
Using this command may cause incompatibility of the QubeCL with the
Control
Software or with some of its modules, so it has to be executed only under assistance
by ppqSense staff.
11.2
Adv.
2 tab
This tab holds fewer controls,
being available for future development.
The tab holds two
panels dedicated to some advanced featurer both of the DDS sub-module and the LIA module.
49 - 69

---
Figure 23:
The first Advanced tab of the Control
Software GUI
The Clock & Sync.
panel
holds a series of advanced controls for the DDS sub-module.
The first three buttons, DDS, allow the user to switch on and off the timing clocks generated
by the DDS board.
The first clock is the one used to serve as a time reference for the DDS
circuits.
If this DDS is switched off, the onboard DDS won’t be able to source any signal.
The Int.
Sync buttons switches on and off the synchronization signal internal to the QubeCL,
used by the driver to establish a shared time base between the DDS sub-module and the LIA
or PDH modules.
Switching off this signal
will
disrupt the synchronization between the DDS
sub-module and the PDH/LIA modules, preventing any locking procedure to work as described
in this manual.
The Ext.
Sync button switches on and off the SYNC OUT signal
available on the C7
connector of the LH module.
The Ext.
Sync.
Phase allow the user to change the phase of the external synchronization
signal
available on the connector C7 of the LH module, if the DDS sub-module is included in
the QubeCL stack.
The Lock-In Amplifier panel
holds some advanced commands interacting with the LIA
module.
The MIXER button switches on and off the demodulating action of the LIA module.
The 90°SHIFT enables or disables a 90 degrees shift on the demodulating signal of the Lock-
In Amplifier.
The Clock Div.
selector enables the user to set a division factor for the demodulating signal
of the LIA module, ultimately reducing the actual
frequency of the demodulating signal.
By changing the value of
each one of
those controls from the standard ones will
affect the
ability of the QubeCL to perform locking with the LIA modules.
50 - 69

---
12
Interfacing with lasers
The basic version of
the QubeCL has two modules that provides the connections to the
laser:
the Laser Head (LH) Module,
with an SMA connector to source current to the laser,
and the Temperature Controller (TC) Module, with a multipolar connector, used to sense and
stabilize the temperature of the laser.
The vast majority of laser do not have an electrical connection that directly matches the output
connectors of the QubeCL, hence ppqSense has developed a variety of interconnection boards
and housing modules in order to make it easier for the user to connect the laser to its driver
without self-assembled solutions.
12.1
Laser Housing for QubeCL
To make the set-up even smaller, reducing space consumption, complexity and electrical noise
susceptibility,
it is possible to include a laser
Housing Module on top of
a QubeCL.
Such
Module provides the mechanical
anchorage for the laser, acts as an heatsink and provides all
the necessary electrical interconnection with the other modules composing the QubeCL device.
When a laser Housing Module is present,
the LH and TC modules do not have the external
connection to the laser, which are no more necessary, being the laser connected to them inside
the QubeCL itself.
12.1.1
Butterfly Packaged laser Housing
ppqSense provides various different typologies of laser Housing Modules for butterfly-packaged
lasers, in order to guarantee compatibility with a variety of different pins layout and package
dimensions.
Modules are available both with optical-fiber output or with mounting holes to
screw a periscope on the side of the QubeCL.
Please, contact ppqSense to verify the compatibility of our Butterfly Housing with the model
you want to buy or to discuss the realization of a dedicated solution.
12.1.2
HHL Packaged laser Housing
A variety of
different HHL housing modules are available to be mounted on a QubeCL,
developed to ensure compatibility with different manufacturer standards.
The modules come
with the necessary holes to mount periscopes and collimating lenses, if needed.
Please, contact ppqSense to verify the compatibility of our HHL Housing with your laser model,
its power ratings and the heat dissipation requirements.
12.2
Laser Interface Boards
The Interface Boards developed by ppqSense are useful to interface a laser with its QubeCL
driver
without including the laser
in the driver
itself.
Interface Boards have two electrical
connections that matches the outputs of the QubeCL driver, while also presenting a connector
dedicated to match the pins of the laser.
The Interface Boards do not provide any mechanical
anchorage for
the laser
and are only
intended for its electrical
connection with the QubeCL driver.
51 - 69

---
12.2.1
HHL Interface Board
The HHL interface boards is specifically designed to provide electrical
interconnection with
the QubeCL driver
in a very small
board (55 x 35 mm),
while also providing some safety
features for the laser.
Please,
contact ppqSense staff to verify the compatibility of
the HHL Interface Board with
your specific laser.
12.2.2
NanoPlus Cubic Housing Interface Board
In order to connect the QubeCL with lasers embedded in the Cubic Mount from NanoPlus,
we’ve developed an interface board which can be directly screwed to the DB9 connectors of the
Cubic Mount, providing all
the external
connection to directly attach to the QubeCL driver.
12.3
High Frequency modulation with Bias-Tee
Some of the Interface Boards and Housing Modules may include an high-frequency modu-
lation input to source the laser with modulating signals at higher frequency with respect to
the ones allowed by the QubeCL modulators (see chapter 6,
table 4 and table 5,
and the
QubeCL datasheet).
To allow such high-frequency modulation,
the Interface Boards and Housing Modules use a
Bias-Tee.
Bias-Tee is an electronic component used to inject high-frequency modulating signal
over
a DC bias signal.
The component has three ports of which one is used to source a DC bias
signal while the other injects the radio frequency signal, the third port is the output one, where
the bias and the RF signal
are superimposed.
The Bias-Tee works to decouple the sources of
the two input signals.
Figure 24:
Bias Tee simplified electronic diagram
As it is shown in figure 24,
Bias Tee can be described as a LC circuit working as a filter to
decouple DC bias and RF signals on their respective ports.
Because of its circuit model, Bias
Tees can only couple RF signal
over a certain cut-off frequency.
The Bias-Tee used inside ppqSense devices has a 500 kHz cut-off frequency, hence it can’t be
used to inject modulation signal
below such frequency.
12.3.1
Bias-Tee usage for high freuqency laser modulation
QubeCLs drivers are able to source modulating signals to the laser by themselves with a -
3 dB bandwidth of
2 MHz,
without
needing any external
circuitry.
The great
advantage
52 - 69

---
of
the modulating signals of
the QubeCL is that they are generated as current signals by
dedicated current-drivers.
The fact that the QubeCLs houses current drivers for modulating
signal
generation makes the process of generating a modulating signal
for the laser an easy
and straightforward operation,
since the amplitude and frequency of the signal
can be easily
defined by the input signal
to the QubeCL or with the DDS interface on the GUI.
Bias-Tee
steps in if higher frequencies modulating signals are necessary.
All
the Housing Modules
and Interface Boards equipped with a Bias-Tee have a high-
frequency modulation input to which the user has to connect the modulating signal.
Bias-Tee has the advantage of withstanding higher frequencies, but the disadvantage of being
designed to couple voltage-signal
and not current-signal
as the one that should be used to
properly and easily drive a laser.
Given that,
when sourcing a signal
to the RF port of
the
Bias-Tee boards, care must be taken in order to properly understand how the voltage applied
to the port is converted into current on the laser.
In those condition, the amplitude of the modulating current not only depends on the ampli-
tude of the modulating signal
but also on the operating point of the laser itself (which is set
by the DC bias current) and on the dynamic resistance that the laser shows at such operating
point.
It is possible to calculate the amplitude of the modulating current using a circuit model
for the Bias-Tee analogous to the one pictured in figure 24.
The circuit model
for the Bias Tee includes a capacitor to isolate the RF port from the DC
bias sourced on the DC port,
while the inductance on the DC port isolates it from the RF
signal
sourced on the RF input.
The aim is to determine the amplitude of
the Modulating
current-signal
that is sourced out of the Bias-Tee RF+DC port knowing the amplitude of the
voltage-signal
applied to the RF port of the device.
For the purpose of calculating the RF signal
amplitude, the inductance on the DC port can
be considered as an open circuit and hence ignored.
We obtain an R-C series circuit, in which
C is the capacitance of the RF port of the Bias-Tee and R is the Dynamic Resistance of the
laser at the specific working condition, which depends on the DC bias current.
Concerning the Capacitor C, the current that flows through it can be calculated as:
dv
i = C
dt
For a Square Wave signal,
this formula can be used to find out I
,
which is the maximum
M
current that can be injected through the capacitor, by using the subsequent values:
• C is the capacitor of the Bias-Tee circuit, in case of the ppqSese Housing Modules and
Interface Boards we have C = 1 µF;
• The derivative of the voltage signal
can be approximated with the voltage peak-to-
peak amplitude of the signal
divided by its rise (or fall) time.
This calculation returns I
, the peak injected current for a laser Dynamic Resistance of 0 Ω.
M
If the Dynamic Resistance is higher, it will
reduce the amplitude of the modulating current.
If
53 - 69

---
the Dynamic Resistance is much higher than the impedance of the Capacitor it is dominant
and can limit the amplitude of the modulating current to a maximum value of:
V
pp
i
=
mod
R
Where V
is the peak-to-peak amplitude of
the modulating signal
and R is the Dynamic
pp
Resistance of the laser.
The two conditions described above are the extreme ones:
• In one case the Dynamic Resistance R of the laser is extremely low and the amplitude of
the modulating signal
is only determined by the value of the Capacitor of the Bias-Tee
and by the characteristics of the modulating signal.
• In the second case,
the Dynamic Resistance R is higher
and predominates over
the
impedance of
the Capacitor at a level
that the impedance can be ignored,
ultimately
limiting the peak-to-peak value of the resulting modulating current.
Depending on the value of the resistance, the system can show one of these two limit behaviours
or something in the middle between the two limit cases if none of the impedances dominates
over the other one.
For a practical approach to the use of the Bias-Tee, once the parameters of the modulating
signal
are set (peak-to-peak amplitude and rise/fall
time), the current I
can be calculated in
M
order to find out its magnitude.
The modulating signal parameters can be consequently modified in order to lower the amplitude
of
I
if
it comes out to be too high.
I
is the overall
maximum amplitude of
the current
M
M
that the bias tee can inject into the laser in any condition,
since its limited by the Bias-Tee
characteristics themselves.
Be careful:
the fact that this is the maximum current that can be injected by the Bias-Tee
does not mean that the laser can withstand it.
This value is the limit for the Bias-Tee but
may be higher than the maximum value allowed for the laser.
If possible,
it could be helpful
to obtain an estimate of the dynamic resistance of the laser,
in order to have a better model
for the circuit and not have to rely solely on the maximum
current value obtained with the previous calculations.
In any condition,
starting with a low amplitude modulating signal,
lower than 100 mVpp,
can be helpful to avoid the risk of damaging the laser.
From this starting point, the amplitude
can be gradually increased trying to find the best amplitude for the modulating signal, always
being careful
not to obtain a too high modulating current.
A series resistor can be added between the signal
generator and the Bias-Tee RF input in
order to clip the amplitude of the modulating signal, obtaining the same effect that would be
obtained with a high-resistance laser that has been described above.
54 - 69

---
12.3.2
Interface Boards Wiring
All
the ppqSense Interface Boards present the three port configuration of the Bias-Tee,
if
the Bias-Tee option is included,
being the DC Bias Current input,
the RF Modulation Input
and the RF + DC Output which is embedded into the laser connector.
Moreover,
since the
Boards acts as a complete interface to the laser,
they also have a TEC connector to allow
Temperature Stabilization.
In figure 25 and figure 26,
the position of the described connection into the two interface
board is displayed.
Figure 25:
Drawings of the NanoPlus Cubic Mount Interface Board
Figure 26:
Drawings of the HHL Interface Board
55 - 69

---
13
Direct Communication
It is possible to communicate with the QubeCL without the QubeCL Control
Software.
This option can be useful
if the user wants to integrate the QubeCL into a complex system
which is managed by a control
software.
The integration can be done by allowing the third
party control
software to directly communicate with the Serial
Port of the QubeCL using the
appropriate commands.
Here we provide all
the technical
information to establish the Serial
connection and a list of
the available commands, along with their syntax and reply from the QubeCL.
CAUTION
The QubeCL Control Software provided by ppqSense provides some safety features that
are useful
to prevent any usage of the QubeCL that may harm both the driver or the
laser connected to it.
While using a third party control
software,
the user must be particularly cautious in
order to avoid any possibly harmful
condition for the QubeCL driver or the laser.
Before using any third party control
software,
ppqSense strongly encourages the user
to contact the ppqSense technical
support to obaint counseling and advices from the
techincal
staff.
13.1
QubeCL Serial
Port Specifications
The QubeCL Control
Software communicates with the QubeCL by the mean of
a Serial
Connection that can also be used by a third party software to interact with the driver.
The
baud rate of the QubeCL Serial
Port is 115200 bps,
which has to be set on the third party
software in order for it to work with the driver.
The connection is possible only if the QubeCL Control Program from ppqSense is not running,
otherwise the Serial
Port resources is kept occupied and it is not accessible from the third
party software.
13.2
QubeCL Serial
Commands
User can find all
the available commands listed in the subsequent tables.
Those commands
are the ones that are available for
the final
user
to be integrated in a third party control
software or to be used in the QubeCL Control
Software to directly communicate with the
driver via the dedicated form in the Advanced tab.
In order for the commands to work, they must be sent at the appropriate Baud Rate and each
command must be terminated with the \n character.
The strings that the QubeCL sends back as a reply, if any for the specific command sent, are
terminated with the \r\n characters.
56 - 69

---
Table 13:
Available commands for QubeCL current generator.
Output
Cmd
Value
R/W Action
Unit
Format
Output Current and Modulations
Returns the identificative code of the
QubeCL-
id:
?
R
N.A.
instrument
####
Returns
the
Output
Current
value
ilas:
?
R
####.##
mA
readed from the QubeCL monitor.
?
R
Returns the Current setpoint.
####.##
mA
iset:
####.##
W
Sets the Current setpoint.
N.A.
mA
on
W
Switches the Current ON.
N.A.
N.A.
iout:
Switches
the
Current
OFF.
If
the
off
W
modulations were active,
they’re also
N.A.
N.A.
switched OFF.
Returns
the
Output
Current
Maxi-
?
R
####.##
mA
imax:
mum Limit set by the user.
Sets
the Output
Current
Maximum
####.##
W
####.##
mA
Limit.
Returns the measured Voltage Drop
vlas:
?
R
####.##
V
across the laser
Enables
the
overall
Modulation
Switch.
This
command
becomes
on
W
N.A.
N.A.
mod:
active 10 seconds after.
the use of
the iout:on command.
Disables
the
overall
Modulation
off
W
Switch.
If
some of
the modulations
N.A.
N.A.
are active, they are switched off too.
on
W
Enables the Modulation Channel
1.
N.A.
N.A.
mod1:
off
W
Disables the Modulation Channel
1.
N.A.
N.A.
on
W
Enables the Modulation Channel
2.
N.A.
N.A.
mod2:
off
W
Disables the Modulation Channel
2.
N.A.
N.A.
57 - 69

---
Table 14:
Available commands for QubeCL temperature controller.
Output
Cmd
Value
R/W Action
Unit
Format
Laser Temperature Controller
Return the Temperature of
the laser
tlas
?
R
####.##
°C
measured by the TC Module.
Switches the Temperature Stabiliza-
on
W
N.A.
N.A.
tstab:
tion on.
Switches the Temperature Stabiliza-
off
W
N.A.
N.A.
tion off.
Return the
laser
Temperature
Set-
?
R
####.##
°C
tset:
point
####.##
W
Sets the laser Temperature Setpoint
N.A.
°C
Sets the Proportional Gain Coefficient
kp:
####.##
W
for the digital
PID that manages the
N.A.
A/K
laser Temperature Stabilization.
Sets the Integral
Gain Coefficient for
ki:
####.##
W
the digital PID that manages the laser
N.A.
A/Ks
Temperature Stabilization.
Sets the Derivative Gain Coefficient
kd:
####.##
W
for the digital
PID that manages the
N.A.
As/K
laser Temperature Stabilization.
Returns the Gain Values of the Digital
PID that manages the laser Temper-
ature Stabilization.
pid:
?
R
kp
####.##:
A/K
ki
####.##:
A/Ks
kd
####.##
As/K
Returns the sign of the PID that man-
?
R
ages the laser Temperature Stabiliza-
#
bool
tecsign:
tion.
0:
reverse, 1:
direct.
Sets the sign of
the digital
PID that
dir, rev
W
manages the laser Temperature Sta-
N.A.
N.A.
bilization.
Returns
the Maximum Temperature
?
R
####.##
°C
tlimax:
Setpoint for the Stabilization PID.
Sets the Maximum Temperature Set-
####.##
W
N.A.
°C
point for the Stabilization PID.
Returns
the
Minimum Temperature
?
R
####.##
°C
tlimin:
Setpoint for the Stabilization PID.
Sets the Minimum Temperature Set-
####.##
W
N.A.
°C
point for the Stabilization PID.
58 - 69

---
Returns the Maximum Current Out-
?
R
####.##
A
teclim:
put for the TC Module.
Sets
the Maximum Current
Output
####.##
W
N.A.
A
for the TC Module.
Returns the threshold for the Active
?
R
####.##
°C*s
teslim:
Temperature Error Control.
Sets the threshold for the Active Tem-
####
W
N.A.
°C*s
perature Error Control.
59 - 69

---
Table 15:
Available commands for the DDS submodule.
DDS Module
Output
Cmd
Value
R/W Action
Unit
Format
Returns
the status
of
the DDS on
modulation channel
1.
?
R
#
Bool.
1:
DDS is active.
dds1:
0:
DDS is inactive.
on
W
Activates the DDS on channel
1.
N.A.
N.A.
off
W
Switches off the DDS on channel
1.
N.A.
N.A.
Returns the waveform set for DDS on
channel
1.
?
R
#
Int.
dds1w:
1:
Sine wave.
2:
Triangular wave.
Sets
the waveform for
the DDS on
channel
1.
Accepted values are:
#
W
N.A.
Int.
1:
Sine wave.
2:
Triangular wave.
Returns the frequency of the DDS on
?
R
####.##
Hz
dds1f:
channel
1.
Sets
the frequency for
the DDS on
####.##
W
N.A.
Hz
channel
1.
Returns the amplitude of the DDS on
?
R
####.##
mA
dds1a:
channel
1.
Sets the amplitude for
the DDS on
####.##
W
N.A.
mA
channel
1.
Returns
the
phase
of
the
DDS on
?
R
####.##
Deg.
dds1p:
channel
1.
Sets the phase for the DDS on chan-
####.##
W
N.A.
Deg.
nel
1.
Returns
the status
of
the DDS on
modulation channel
2.
?
R
#
Bool.
1:
DDS is active.
dds2:
0:
DDS is inactive.
on
W
Activates the DDS on channel
2.
N.A.
N.A.
off
W
Switches off the DDS on channel
2.
N.A.
N.A.
Returns the waveform set for DDS on
channel
2.
?
R
#
Int.
dds2w:
1:
Sine wave.
2:
Triangular wave.
Sets
the waveform for
the DDS on
channel
2.
Accepted values are:
#
W
N.A.
Int.
1:
Sine wave.
2:
Triangular wave.
60 - 69

---
Returns the frequency of the DDS on
?
R
####.##
Hz
dds2f:
channel
2.
Sets
the frequency for
the DDS on
####.##
W
N.A.
Hz
channel
2.
Returns the amplitude of the DDS on
?
R
####.##
mA
dds2a:
channel
2.
Sets the amplitude for
the DDS on
####.##
W
N.A.
mA
channel
2.
Returns
the
phase
of
the
DDS on
?
R
####.##
Deg.
dds2p:
channel
2.
Sets the phase for the DDS on chan-
####.##
W
N.A.
Deg.
nel
2.
Returns
the frequency of
the
DDS
?
R
channel
to which the Sync.
Out sig-
####.##
Hz
syncf:
nal
is synchronized.
Synchronizes the Sync.
Out signal
to
ch1
W
N.A.
N.A.
the DDS1
Synchronizes the Sync.
Out signal
to
ch2
W
N.A.
N.A.
the DDS2
61 - 69

---
Table 16:
Available commands for the PLL locking module.
PLL Locking Module
Output
Cmd
Value
R/W Action
Unit
Format
Allow the user to select which signal
is present on the DIVIDER MON out-
put.
Allowed values are:
mux:
#
W
N.A.
Int.
0 DIVIDER MON switched off.
2 DIVIDER MON set on RF.
4 DIVIDER MON set on LO.
Set
the sign of
the PLL correction
loop.
Accepted values are:
sig:
#
W
N.A.
Int.
0:
negative correction loop.
1:
positive correction loop.
off
W
Switches off the PLL.
N.A.
N.A.
pll:
on
W
Switches on the PLL.
N.A.
N.A.
Returns the value of cp, which is the
?
R
gain of
the PLL.
Must
be switched
##
HEX.
ON.
cp:
Switches ON the proportional
gain of
on
W
N.A.
N.A.
the PLL chip.
Switches OFF the proportional gain of
off
W
N.A.
N.A.
the PLL chip.
Set the proportional
gain.
The value
#
W
N.A.
Int.
can range from 1 to 8.
Sets the divisor coefficient for the RF
ndiv:
####
W
N.A.
Int.
frequency input.
Sets the divisor coefficient for the LO
rdiv:
####
W
N.A.
Int.
frequency input.
Sets the PLL correction loop in pro-
off
W
N.A.
N.A.
pby:
portional
only mode.
Sets
the
PLL
correction
loop
in
on
W
N.A.
N.A.
proportional-integrative mode.
Sets the Integrative time constant of
tp:
#
W
the PLL correction loop.
Accepted
N.A.
Int.
values range from 0 to 3.
Sets
the
Proportional
gain
of
the
tz:
#
W
PLL correction loop.
Accepted val-
N.A.
Int.
ues range from 0 to 3.
Sets the proportional
gain of the PLL
hg:
#
W
chip.
Accepted values ranges from 0
N.A.
Int.
to 3.
62 - 69

---
Switches OFF the correction current
off
W
injected by the PLL module into the
N.A.
N.A.
lk:
Bias Current.
Activates
the correction current
in-
jected by the PLL module into the
on
W
N.A.
N.A.
Bias
Current.
Overall
modulation
must be enabled for it to work.
Returns the value of the PLL correc-
lm:
?
R
####.##
mV
tion loop output voltage.
63 - 69

---
Table 17:
Available commands for the PDH locking module.
PDH Locking Module
Output
Cmd
Value
R/W Action
Unit
Format
Sets the PDH correction loop in pro-
off
W
N.A.
N.A.
pdhint:
portional
only mode.
Sets
the
PDH correction
loop
in
on
W
N.A.
N.A.
proportional-integrative mode
Deactivates the HOLD input of
the
off
W
N.A.
N.A.
pdhhold:
PDH module.
Activates
the HOLD inptut
of
the
on
W
N.A.
N.A.
PDH module.
Deactivates the correction current in-
off
W
jected into the Bias
current
by the
N.A.
N.A.
pdhlock:
PDH module.
Activates
the correction current
in-
jected into the Bias
current
by the
on
W
N.A.
N.A.
PDH module.
Overall modulation en-
able must be active for it to work.
Sets the first proportional
gain of the
pdhrint:
#
W
PDH correction loop.
Accepted val-
N.A.
Int.
ues range from 0 to 3.
Sets the second proportional
gain of
the PDH correction loop.
Also affect
pdhtz:
#
W
N.A.
Int.
the Integrative time constant.
Ac-
cepted values range from 0 to 3.
Sets the time constant
of
the PDH
pdhtp:
#
W
correction
loop.
Accepted
values
N.A.
Int.
range from 0 to 3.
Sets
which signal
is
present
on the
Monitor Output of the PDH module.
pdhmon:
#
W
Accepted values are:
N.A.
Int.
0:
Error Signal.
1:
Correction signal.
Returns the values of
the Error
and
####.##:
pdhmonint: ?
R
Correction signals converted by the on
mV
####.##
board ADCs.
Returns the input offset of
the PDH
?
R
####.##
mV
pdhvoff:
correction loop.
Sets the input offset of the PDH cor-
####.##
W
rection loop.
The value can range
N.A.
mV
from 0 to 5000.
Default is 2500.
64 - 69

---
Returns the PDH correction loop pre-
?
R
####.##
mV
pdhdp:
amplifier gain.
Sets
the PDH coorection loop pre-
####.##
W
amplifier gain.
Accepted values range
N.A.
mV
from 0 to 63.
65 - 69

---
Table 18:
Available commands for the LIA locking module.
LIA Module
Output
Cmd
Value
R/W Action
Unit
Format
Returns the mode of
the correction
loop:
?
R
#
Bool.
0 stands for P mode.
lkpi:
1 stands for PI mode.
0
W
Sets the correction loop in P mode.
N.A.
Int.
1
W
Sets the correction loop in PI mode.
N.A.
Int.
Returns the status of
the Low Pass
Filter of the correction loop:
?
R
#
Bool.
0 stands for inactive.
lkflt:
1 stands for active.
en
W
Activates Low Pass Filter
N.A.
N.A.
dis
W
Deactivates Low Pass Filter
N.A.
N.A.
?
R
Returns the VGA Gain.
####.##
dB
lkgain:
####.##
W
Sets the VGA Gain
N.A.
dB
Returns
the
pole
of
the
correction
?
R
#
Int.
lktp:
loop.
Sets the pole of
the correction loop.
#
W
N.A.
Int.
Accepted value ranges from 0 to 3.
Returns the proportional
gain of
the
?
R
#
Int.
lktz:
correction loop.
Sets the proportional
gain of the cor-
#
W
rection loop.
Accepted values ranges
N.A.
Hz
from 0 to 3.
Returns the frequency setting for the
?
R
#
Int.
lktpb:
Low Pass Filter of the correction loop.
Sets the frequency setting for the Low
#
W
Pass Filter of the correction loop.
Ac-
N.A.
mA
cepted value ranges from 0 to 3.
on
W
Activates the LIA correction loop.
N.A.
N.A.
lklock:
off
W
Deactivates the LIA correction loop.
N.A.
N.A.
Returns
the demodulation mode of
the LIA.
?
R
0 stands for f mod.
#
Int.
1 stands for 2f mode.
lkdemod:
2 stands for no deomdulation acrive.
f
W
Sets the f demodulation mode
N.A.
N.A.
2f
W
Sets the 2f demodulation mode
N.A.
N.A.
free
W
Deactivates the demodulation
N.A.
N.A.
66 - 69

---
Returns the currently set for the out-
put monitor of the LIA module:
?
R
####.##
Deg.
0 stands for correction monitor.
lkmon:
1 stands for ERROR monitor.
Sets the output monitor to show the
err
W
N.A.
N.A.
ERROR signal.
Sets the output monitor to show the
lock
W
N.A.
N.A.
correction signal.
Returns the digital
filter currently in
use:
1 stands for BP0.
2 stands for BP1.
?
R
3 stands for BP2.
#
Int.
4 stands for LP1.
5 stands for LP2.
6 stands for Notch.
lkIIR:
7 stands for All
Pass.
BP0
W
Sets the BP0 filter
N.A.
N.A.
BP1
W
Sets the BP1 filter
N.A.
N.A.
BP2
W
Sets the BP2 filter
N.A.
N.A.
LP0
W
Sets the LP0LP0 filter
N.A.
N.A.
LP1
W
Sets the LP1 filter
N.A.
N.A.
NOTCH
W
Sets the NOTCH filter
N.A.
N.A.
ALLPASS
W
Sets the ALLPASS filter
N.A.
N.A.
Table 19:
Available commands to check the instrument status.
QubeCL Instrument Status
Returns
the
Supply
Voltage
for
vcc:
?
R
the
CM Module
measured
by
the
####.##
V
QubeCL.
Returns the internal
Temperature of
tsense:
?
R
####.##
°C
the QubeCL instrument.
67 - 69

---
14
Customer support and useful
contacts
The ppqSense staff can be reached for
any problem or
request concerning the QubeCL
devices.
To address any malfunctions or request help while working with our QubeCL, please write an
email to qube.support@ppqsense.odoo.com, our staff will reply to the raised ticket as soon
as possible.
To obtain quotations, support or updates concerning the purchase of our instruments, please
write an email
to info@ppqsense.com.
All
the available documentation concerning both the QubeCL and all
our other products
can be found on our website:
www.ppqsense.com.
ppqSense s.r.l.
Viale L. Ariosto 492/B
50019, Sesto Fiorentino (FI)
Italy
+39 055 80 23 943
68 - 69

---
Revision History
Revision
Date
Author(s)
Description
1.0
2024.05.15
LM
First redaction.
69 - 69